# Immersive UV Visual Stimulus Platform

Jay Borseth 2013.12.04

## 24” Dome, No Mirrors

- Dome diameter : 24”

- Eye to dome: 11.2 cm

- 240 degree stimulation

- Dome to lens:  39.8” assuming LightCrafter 4500

## 24” Dome, single mirror

Folded light path:  29” deep (green)  x 24”high (blue) x 24” wide

## Multiple mirrors with 12” / 18” dome

### Projection Issues

Will the image be in focus over the entire surface since the light path lengths differ? 
General Characteristics

## Two eye stimulus

- 240 degree visual stimulus

- 270 degree visual stimulus

## Spatial resolution

- 6 pixels per degree at equator  of dome 
1080x1080  @180 degrees

- 5 pixels per degree at equator  of dome 
LCr4500 912x912 @180 degrees

- 3.3 pixels per degree at equator of dome 
LCr3000 608x684 @180 degree

## Bits per pixel

- 8 bits (implies higher end DLP chips)

- 7 bits (lower end DLP chips)

## Mouse platform

- Air bearing sphere

- Wheel

## Mouse eye sensitivity

## Visual Channels

C. palustris  Panasonic GH-2, Coastal Optics 60 mm f/4 APO lens, daylight.

- One channel (UV & green together)

- Two channels (UV & green separate)

- Three channels (UV, blue, green)

- If only one channel, what is the relative proportion of UV/Green?

- If color source imagery is used, what is the mapping to monochrome?
Y = 0.2126 R + 0.7152 G + 0.0722 B (NTSC)

## Typical Dome plastics

MC is the most common Plexiglas.

http://www.cleardome.com/domes.htm uses a comparable plastic called Acrylite FF which looks like:

Solid line upwards at 340nm is Acrylite FF

## Light source

- Broad spectrum stimulus (Xenon lamp)

- Narrow spectrum (LEDs)

- Brightness range:  maximum brightness of 20 mCd/m2

# Stimulus

## Framerate

- 60 Hz

- 120 Hz

- 180 Hz

## Movie Playback

- 60 Hz

- 120 Hz

- 180 Hz

## Patterns and Stimulus

- Grayscale Gradients

- Color Gradients / Objects

- Object scaling size to simulate depth

- Object occlusion

- Maximum number of objects displayed simultaneously

- 3D

# DLP

## Chipsets and Maximum Pattern Rates

### DLP and UV

Problems with UV and DLP devices

- In many non-UV designed system there is a plastic lens called the “fly eye” in the light path which will yellow after just a day of high UV exposure.

- Gas used between the cover and DLP mirrors acts as a lubricant for DLP mirrors and “thickens” with UV exposure.  When it gets thick, the mirrors stop moving.

## Hardware Options

### LC3000-PRO DLP Pico Projector with UV LED DLP Pico

$1450

http://keynotephotonics.3dcartstores.com/LC3000-PRO-DLP-Pico-Projector-with-UV-LED-DLP-Pico-Projector_p_48.html

- 0.3” DLP

- ProjectorThrow ratio = 1.66

- Focus range 364mm to 2169mm (85.39 inches)

- width = distance / throw ratio

- 12” dome, dome to lens is 19”

- 18” dome, dome to lens is 29.88”

- 24” dome, dome to lens is 39.84”

#### Issues

- 120 Hz @ 8-bits, 180 Hz @ 7-bits  CONFIRM!!!

### LightCrafter 4500

- 0.45” DLP

- Throw ratio = 1.66

- Focus range 300mm to infinity

- width = distance / throw ratio

- 12” dome, dome to lens is 19”

- 18” dome, dome to lens is 29.88”

- 24” dome, dome to lens is 39.84”

#### Issues

- As of Dec 3, 2013, unit shuts down with LEDs removed.  Can’t substitute light modules. Firmware update promised.

- 120 Hz @ 8-bits, 180 Hz @ 7-bits

- Must hack each application to supply time series monochrome as RGB

- Not rated for UV.  “Fisheye” lens will yellow with UV exposure.

### LightCrafter Pro4500 for UV

http://www.wintechdigital.com/product_s.asp?id=17

~$2000

- Manufacturer of LCr sold by TI

- New product, available end of Dec 2013

- Pro model handles UV includes glass replacement for “flyeye” lens and different coating for the 20 other lenses in the light module.

#### Issues

- Same as regular LCr4500

- Not yet available

### CEL5500-UV

https://www.dlinnovations.com/wp/?page_id=759

~$6500

- .55” DLP

- UV specific model available

- 37mm projection lens ring

#### Issues

X3 Machine Vision Solution with DLP 5500 (X3-PM55)

http://keynotephotonics.3dcartstores.com/X3-Machine-Vision-Solution-with-DLP-5500-X3-PM55_p_14.html

$3999

- .55” DLP

- No optics

- Can be converted to HDMI 8bit@180Hz monochrome for ~40K firmware development

#### Issues

### VPixx ProPixx

http://www.vpixx.com/products/visual-stimulus-displays/propixx-lite.html

$31K

- 180 Hz native projection

- Run any software application without modification at 180Hz

#### Issues

- $31K

# Software

## Operating System

- Windows

- Mac

- Linux

## Applications

- PsychoPy

- matlab

- visionegg

- other